Central Hub — Setup Guide

MOXA / LAN Instrument Integration — Radwag Balance & LabIndia pH Meter

LIMS Instrument Integration Collector  ·  balance_server.py

The Central Hub (balance_server.py) is the collector service that captures printouts from Radwag balances and LabIndia pH meters — either through a MOXA NPort serial-to-Ethernet gateway or directly over the instrument's LAN port — and writes the records to Supabase, where the LIMS reads them. It runs as a small local web app on http://127.0.0.1:8000 (Balance / pH Meter / Settings tabs).

1.  Download the repository from GitHub

Get the project from GitHub: https://github.com/arupa444/moxaEquipLIMS

Download ZIP:

Open “https://github.com/arupa444/moxaEquipLIMS” in a browser.

Click the green Code button → Download ZIP.

Right-click the ZIP → Extract All… → choose a location.

Place the folder somewhere permanent, e.g. C:\LIMSBalanceHub.

2.  Prerequisites (once per machine)

Windows 10 or 11, 64-bit (Windows Server 2016+ also works).

Python 3.11 x64 — download from https://www.python.org/downloads/ and tick “Add python.exe to PATH” during install (3.10 is the minimum; 3.11 recommended).

Network: the hub PC must reach the instruments / MOXA (same LAN/subnet) and reach Supabase (self-hosted API gateway, e.g. http://10.1.11.98:8000, or the cloud project).

The Supabase database tables must already exist (apply the migrations — Step 5).

3.  Start the hub (one click — it sets everything up)

Double-click start-hub.bat in the repo folder. On the first run it automatically:

creates the Python virtual environment (.venv),

installs all required packages (requirements.txt),

starts the hub hidden in the background, serving http://127.0.0.1:8000,

writes logs to hub.out.log and hub.err.log.

On later runs it just starts the hub (packages reinstall only when requirements.txt changes).

To stop the hub, double-click stop-hub.bat. The hub refuses to start a second copy while one is running.

Manual alternative (visible console):

.venv\Scripts\python.exe balance_server.py

4.  Get the admin token and sign in

The hub is protected by a per-install admin token (no password to set), generated on first run and saved to data\admin_token.txt.

1. Read your token:

type C:\LIMSBalanceHub\data\admin_token.txt

2. Open http://127.0.0.1:8000 in a browser.

3. Paste the token into Admin token → Sign in.

The session expires after 20 minutes of inactivity — just paste the token again.

To rotate the token: delete data\admin_token.txt and restart the hub.

To pin a fixed token across machines: set the HUB_ADMIN_TOKEN environment variable.

5.  Configure Supabase (mandatory — nothing is captured until done)

In the hub, go to Settings and fill in the connection. Deployment:

Self-hosted Supabase (via API URL) — the usual choice. Connects through the REST API at SUPABASE_URL (e.g. http://10.1.11.98:8000) with the service-role key.

Cloud Supabase — same, pointed at your *.supabase.co project.

| Field | Value |
| --- | --- |
| SUPABASE_URL | your API gateway URL, e.g. http://10.1.11.98:8000 |
| Service role key | the Supabase service-role key |
| Publishable / anon key | the anon key |
| Session secret | any random string, minimum 32 characters |

Click Save & connect. Secrets are stored encrypted at rest (Windows DPAPI) in the git-ignored .cred file; they display only as a set badge afterward.

The Postgres host/port/db/user/password fields are only a fallback used when no SUPABASE_URL is set — normally leave them and rely on the URL.

6.  Add the equipment (MOXA gateway or direct-LAN instrument)

On the Balance (or pH Meter) tab, use Add equipment:

| Field | What to enter |
| --- | --- |
| IP | the device's IP — the MOXA NPort's IP (e.g. 192.168.127.150) or the instrument's own LAN IP |
| Port | the data port — 4001 for a MOXA NPort serial port 1 (400X for port X), or the instrument's TCP-server port |
| Instrument type | Balance or pH Meter |
| Name of the Device | leave blank to auto-fetch the MOXA server name via SNMP, or type a name |

Device-side requirements:

MOXA NPort: Operating Settings → serial port Operation mode = TCP Server, Local TCP Port = 4001. Serial params match the instrument (Radwag ≈ 9600 8-N-1).

Direct-LAN instrument: must run in TCP Server mode on that port and route its printout to the Ethernet channel (not to a separate network-printer address).

The gateway row shows a live Status (auto-refreshes every second):

Connected — link is up.

Disconnected — “Not Connected to the Server. Please plug in the Ethernet cable.” (device unreachable).

Not ready — reachable but the port is busy / not in TCP-Server mode (clears itself; keeps retrying).

The hub auto-reconnects on cable/power loss — no need to re-add the device or restart the MOXA.

7.  Verify end to end

On the balance: print a session — Header → Weight(s) → Footer. On the pH meter: print a Calibration/Readings/Verification set, then a Sample.

The gateway status detail changes from idle → saved ….

In the LIMS (Balance Integration / PH Meter Integration pages), click Refresh — the record appears under Department → Instrument ID → Weight / Adjustment (balance) or Calibration & Verification / Sample (pH), with Operator, Reg No and auto Sample Name.

8.  Run automatically at boot (optional)

To keep the hub running across reboots, register start-hub.bat with Task Scheduler:

Trigger: At startup (or At log on).

Action: Start a program → the full path to start-hub.bat.

Run whether the user is logged on or not.

9.  Troubleshooting

| Symptom | Cause / Fix |
| --- | --- |
| “Supabase configuration is required” | Fill in Settings (Step 6). |
| Gateway shows “Disconnected” | Device unreachable. ping <device IP> from the hub PC; check cable/link LED, subnet, and that the device IP is not the same as the PC's IP. |
| “Not ready / port 4001 busy” | Device reachable but not in TCP-Server mode, or holding a stale session. Set the NPort/instrument to TCP Server; it clears on its own (or set NPort Max Connection = 2). |
| “bad host” when adding | The IP has an invalid character (space/backtick). Re-type a clean IP like 192.168.1.100. |
| Connected but no records after Print | The instrument's print output is not routed to the LAN channel the hub reads (e.g. Radwag Printer set to a network-printer address). Point the print at the hub / the server port the hub uses. |
| Can't reach http://127.0.0.1:8000 | The hub isn't running — double-click start-hub.bat; check hub.err.log. |
| See raw bytes a device sends | Run python capture_lan.py 9100 on the hub, point the device's print output at the PC's IP:9100, press Print. |

10.  Security notes

Access is gated by the admin token (data\admin_token.txt); the dashboard idle-expires after 20 minutes.

Supabase secrets are DPAPI-encrypted at rest in .cred — never commit .cred or the data\ folder (both git-ignored).

The instrument Ethernet link carries instrument data only — it is not used for internet access.

Removing a gateway requires the admin token as confirmation.

Hub URL: http://127.0.0.1:8000   ·   Start: start-hub.bat   ·   Stop: stop-hub.bat   ·   Token: data\admin_token.txt